Section V: Schedule of Requirements

6.1 List of Goods & Delivery Schedule

Unit-: Head Office, Ocean University of Sri Lanka, Mattakkuliya.

| Line-Item No | Description of Goods | Final Destination | Delivery Date |  |
| --- | --- | --- | --- | --- |
|  |  |  | Purchaser’s Allowed Delivery Date | Bidder’s Offered Delivery Date |
|  | - Fume Hood. - Top Loading Balance (Professional Level). - Portable DO meter. - Hot Air Oven. - Water Bath. - Binocular Microscope. - Stereo Zoom Microscope. - Autoclave. - Mini Centrifuge (Refrigerated). - Desiccator. - Counting chambers Sedgewick Rafter. - Surface Plankton Net. - Micro Pipette sets (Ranges: 0.5–10 µL, 10–100 µL, and 100–1000 µL). - Scientific Calculator. - Safety Goggles. - Burette/Funnel Stand. - Spirit lamps. - Tripod Stand. - Crucible Tongs. - Crucible with lid. - Forceps. - Chambers CE Neubauer-imp. w/out. - Gauge Wire Iron packets. - Vortex Mixer. - Desiccator (300mm). - Blender. - Hand sealer. - Mini Biological Safety Cabinet. - Gass Cylinder. - Electric Pressure Cooker 6L. - Spectrophotometer. - Secchi disk. - Vacuum Filtration unit. - Colorimeter. - Digital stirring hot plate. - Grab sampler. - Wooden Herbarium Press. - Herbarium sheets (Pack of 25). - Spectrophotometer. - Spectrophotometer. - Spectrophotometer. - Power Protection (Online UPS) - Buchner Funnel. - Benchtop pH/Ion Meter. - Semi-Automated Kjeldahl Distillation System. - Soxhlet Extraction Apparatus (3-Place unit). - Laboratory Muffle Furnace. - High-Performance Fume Hood. - Rutner sampler. - Heating mantle. - Laboratory Water Distiller. - Top loading balance. - BOD Incubator (Refrigerated). - Muffle furnace. - Thermometer (digital). - Water quality multi meter. - Centrifuge. - Magnetic stirrer. - pH meter with probe. - Dissecting Set (Fish). - Dissecting Tray with Wax. - Thermometer. - Water Quality Testing Meter. - Refractometer. - Analytical Balance. | Main Stores, Crow Island, Mattakkuliya, Colombo – 15. | From the date of Acceptance Maximum 90 days |  |

Name of the Bidder     -:

Address of the Bidder -:

Company Seal -:

Date -:

Section V

6.2 Technical Specification & Compliance - Laboratory Equipment for Regional Centers of the University

| Item No. | Component Description | Minimum Specifications | Quantity | Bidder’s Response (Yes/No) | If “No” comments on the offer |
| --- | --- | --- | --- | --- | --- |
| 1 | Fume Hood | Manual control system; Approx. dimensions: Overall 1750×885×2245 mm, Work chamber 1500×640×1240 mm; Cold rolled steel body with powder coating; Chemical-resistant phenolic resin work surface; 5 mm toughened glass adjustable sash (min opening 690 mm); Air velocity 0.5–0.8 m/s; Noise <65 dB; Light ≥1000 lux; Includes base cabinet (anti-corrosion) or stand; Utilities: sink (PP), gas tap, vacuum/fluid lines; Power: 220–240V, 50/60 Hz with 4 sockets; Alarm system for airflow; Includes external ducting with PP blower & 225 mm PVC duct; Warranty: minimum 24 months (parts & labor); Includes installation, training, calibration, and preventive maintenance during warranty; Spare parts availability for 5 years; Manufacturer authorization & brochures required | 1 |  |  |
|  |  | Manual control system; Approx. dimensions: Overall 1750×885×2245 mm, Work chamber 1500×640×1240 mm; Cold rolled steel body with powder coating; Chemical-resistant phenolic resin work surface; 5 mm toughened glass adjustable sash (min opening 690 mm); Air velocity 0.5–0.8 m/s; Noise <65 dB; Light ≥1000 lux; Includes base cabinet (anti-corrosion) or stand; Utilities: sink (PP), gas tap, vacuum/fluid lines; Power: 220–240V, 50/60 Hz with 4 sockets; Alarm system for airflow; Includes external ducting with PP blower & 225 mm PVC duct; Warranty: minimum 24 months (parts & labour); Includes installation, training, calibration, and preventive maintenance during warranty; Spare parts availability for 5 years; Manufacturer authorization & brochures required |  |  |  |
| 2 | Top Loading Balance (Professional Level) | Capacity ≥220 g; Readability 0.0001 g; Repeatability 0.0001 g; Linearity ±0.2 mg; Min. load 10 mg; Response time ≤2 sec; Electromagnetic force restoration weighing sensor; Auto/self-calibration (temperature/time-based); Stainless steel pan (~90 mm Ø); Multiple weighing modes (standard, counting, %); Interfaces: USB & RS232; GLP/ISO compliant; Backlit display; Operating temp. 18–30°C; Power: 220–230V, 50/60 Hz; Includes installation & documentation; Warranty: minimum 24 months; Spare parts availability; Manufacturer authorization required | 5 |  |  |
| 3 | Portable DO Meter | Portable unit for field & lab use; Optical luminescence DO probe (non-consumptive); Measurement range: 0–20 ppm / 0–200% or better; Accuracy ±0.2 ppm; Resolution 0.01 ppm; Automatic compensation for temperature, salinity, and air pressure; Temp. range 0–50°C; Response time ≤30 sec; Automatic calibration (2-point preferred); Data storage ≥500 readings with date/time; USB connectivity; Backlit LCD display; IP57 or better; Power: battery operated; Includes carrying case, probe, cables, and software; Certification: CE/RoHS; Warranty (meter & probe); Manufacturer authorization required | 1 |  |  |
| 4 | Hot Air Oven |  | 2 |  |  |
|  |  | Capacity ≥40 L; Temperature range 50–200°C; Temperature accuracy ±2°C with PID controller; Inner dimensions approx. 350×350×350 mm; Includes 2 stainless steel shelves; Steel exterior with corrosion-resistant coating; Heater power ~1 kW; Power supply 230V, 50 Hz, single phase; Includes installation; Warranty minimum 12 months; Spare parts availability; Manufacturer/local agent authorization required |  |  |  |
| 4 | Water Bath | Capacity ~20 L; Temperature range ambient +5 to 99.9°C; Resolution 0.1°C; Accuracy ±0.3°C; Uniformity ±0.3°C; PID temperature control with digital display; Stainless steel (SS 304) inner chamber (seamless) and exterior; Heater power ~1 kW; Safety features: dry-run and over-temperature protection; Timer function; Includes lid and accessories; Power supply 230V, 50 Hz; Certification: CE, ISO compliant; Includes installation and documentation; Warranty minimum 12 months; Spare parts availability; Manufacturer/local agent authorization required | 1 |  |  |
| 5 | Binocular Microscope | Brightfield microscope with infinity-corrected optics; Binocular head (360° rotatable, 30° inclined) with interpupillary adjustment (48–75 mm) and diopter adjustment; Eyepieces 10× (FN ≥20, high eyepoint); Magnification range 40×–1000×; Quadruple revolving nosepiece; Objectives: 4×, 10×, 40×, 100× (oil), anti-fungus treated; Coaxial coarse & fine focusing with stop mechanism; Double-layer mechanical stage; Durable metal construction; Includes installation and documentation; Warranty minimum 12 months; Manufacturer/local agent authorization required | 12 |  |  |
| 6 | Stereo Zoom Microscope | Brightfield stereo microscope with Greenough optical system; Binocular head (45° inclined, 360° rotatable) with interpupillary (≈51–75 mm) and diopter adjustment; Eyepieces 10× (FN ~21, high eyepoint); Zoom magnification range ~7×–45×; Working distance ~100 mm; Parfocal achromatic optics, anti-fungus treated; Coarse focusing; Fixed arm stand with stage (≈100 mm) and specimen clips; LED illumination (transmitted) with brightness control; Power supply with rechargeable battery option; Includes installation and documentation; Warranty and spare parts availability; Manufacturer/local agent authorization required | 3 |  |  |
| 7 | Autoclave | Portable Top-loading autoclave for small-load sterilization; Capacity ≤16 L; Stainless steel chamber and lid (seamless construction); Operating temperature ~121°C with timer control; Working pressure ~15 PSI; Digital controller with automatic cycle and timer; Safety features: over-pressure protection and auto shut-off; Efficient sealing with gasket; Power supply 230–240V, 50/60 Hz, single phase; Includes basic accessories, installation, and documentation; Certification: CE compliant; Warranty minimum 12 months; Spare parts availability; Manufacturer/local agent authorization required | 1 |  |  |
| 8 | Minin Centrifuge (Refrigerated) | Capacity up to 18 × 1.5/2.0 mL tubes; Speed range up to ~14,000 rpm; Max RCF ~16,800 × g; Refrigerated with temperature range ~0°C to +40°C; Timer with continuous run function; Low noise operation (<60 dB); Aerosol-tight rotor for safety; Digital control with automatic lid lock/release; Power supply 220–240V, 50/60 Hz; Includes rotor and basic accessories; Includes installation and documentation; Warranty minimum 24 months; Manufacturer/local agent authorization required | 1 |  |  |
| 9 | Desiccator | Non-vacuum desiccator (no stopcock); Made of borosilicate glass (3.3) with high chemical and thermal resistance; Ground flat flange with lid (interchangeable knob); Autoclavable; Suitable for drying and sample storage; Includes standard components (excluding discs); Spare parts availability | 1 |  |  |
| 10 | Counter Chambers Sedgewick Rafter | Chamber dimensions 50 × 20 × 1 mm (1 mL capacity) with 1 mm ruled grid; Made of glass; Supplied with compatible cover glass (~60 × 30 mm); Suitable for plankton/microorganism counting | 3 |  |  |
| 12 | Surface Plankton Net | Standard plankton net with stainless steel ring and conical net; Mesh size ~55 µm (for phytoplankton collection); Synthetic mesh material; Includes removable collection bucket with sieve window and stopcock; Suitable for qualitative surface sampling; Durable construction for field use; Includes basic accessories; Manufacturer authorization required | 2 |  |  |
| 13 | Micropipette sets (Ranges: 0.5-10 µL, and 100-1000 µL | Adjustable volume micropipette with digital volume setting; Each set comprises three pipettes. Pipette Ranges: 0.5–10 µL, 10–100 µL, and 100–1000 µL.; Accuracy ≤±1% and high precision; Fully autoclavable (121°C, 20 min); Chemically resistant construction with corrosion-resistant piston (e.g., stainless steel/PTFE); Compatible with standard tips; Ergonomic single-hand operation with tip ejector; Suitable for precise liquid handling; Includes documentation; Calibration support available; Warranty minimum 12 months; Spare parts availability; Manufacturer/local agent authorization required | 4 |  |  |
| 14 | Scientific Calculators | Non-programmable scientific calculator (equivalent to Casio fx-991CW or similar); ≥600 built-in functions; High-resolution display; Dual power (solar + battery); Supports advanced mathematical and statistical functions; Suitable for academic/examination use; Includes user manual; Warranty; Manufacturer authorization required | 10 |  |  |
| 15 | Safety Goggles | Anti-splash, enclosed, dust-proof safety goggles; Clear lenses with wide field of vision; Comfortable adjustable fit; Complies with EN 166 / CE standards | 15 |  |  |
| 16 | Burette/Funnel Stand | Metal stand with heavy, stable base and vertical rod (~60–90 cm); Adjustable clamp with rubber-coated grip; Suitable for holding standard burettes (25–50 mL); Stable, corrosion-resistant construction | 15 |  |  |
| 17 | Spirit lamps | Made of soda-lime glass; Capacity ~100 mL; Supplied with cap, cotton wick, and holder; Suitable for laboratory heating applications | 10 |  |  |
| 18 | Tripod Stand | Construction: Zinc-plated cast iron top; 15cm Outer Diameter. Dimensions: Approx. 21cm height. | 05 |  |  |
| 19 | Crucible Tongs | Material: Chrome-plated steel. Design: Straight jaws, 15 cm length. | 14 |  |  |
| 20 | Crucible with lid | Material: High-grade porcelain (ceramic), Surface: Glazed (inside & outside except bottom), Capacity: 25 mL, working temperature: 1000–1100 °C, Color: White | 06 |  |  |
| 21 | Forceps | Material: High-quality stainless steel. Design: Straight, pointed tips, length 115 mm. | 20 |  |  |
| 22 | Chambers CE Neubauer-imp. w/out | Design: Improved Neubauer counting chamber; Dark line. Grid: 3 x 3 mm total area with 0.1 mm depth. | 05 |  |  |
| 23 | Gauge Wire Iron packets | Dimensions: 15 cm square. Features: High-temperature ceramic center for uniform heat distribution. | 10 |  |  |
| 24 | Vortex Mixer | Speed range: Adjustable from 0 to 2,500 rpm, Toch or continuous operation,230 V, 50 Hz | 02 |  |  |
| 25 | Desiccator (300mm) | Diameter- 300mm | 01 |  |  |
| 26 | Blender | Mixer Grinder, 3 Stainless Steel Jars, RPM: 20000, Water Drain System | 01 |  |  |
| 27 | Hand sealer | 4-inch sealing length.2mm sealing width | 01 |  |  |
| 28 | Mini Biological Safety Cabinet | Internal work area width -3ft, Class II Type A2 Mini Biological Safety Cabinet, centrifugal fan with automatic speed compensation, H14 HEPA filter, Motorized or manual sliding UV-resistant tempered glass sash, Integrated UV Germicidal System | 02 |  |  |
| 29 | Gass Cylinder | 12.5 Kg | 01 |  |  |
| 30 | Electric Pressure Cooker 6L | Stainless steel outer body,32.5D x 21W x 19.2H Centimeters | 01 |  |  |
| 31 | Spectrophotometer | Visible Range Spectrophotometer, Bandwidth ≤ 6 nm, Wavelength Range 350-1020nm.Digital Lab Tungsten Lamp | 01 |  |  |
| 32 | Secchi disk | 9-inch Secchi Disc with graduated String | 02 |  |  |
| 33 | Vacuum Filtration unit | Complete Vacuum Filtration Kit with Filtering Flask, Hand Vacuum Pump and Buchner Funnel (250 ML), | 01 |  |  |
| 34 | Colorimeter | Portable Colorimeter Color Analyzer, Digital, Measuring Aperture 8mm, Color Difference Meter, High-resolution digital color display (minimum 2-inch) showing color values and difference charts. Rechargeable Lithium-ion battery with a compatible 230V, 50Hz charging adapter.1- 2 Years Comprehensive Warranty. | 01 |  |  |
| 35 | Digital stirring hot plate | Top Plate Material: Ceramic-coated or tempered glass / stainless steel, Plate Size: Approximately 135–180 mm diameter (or equivalent square size), Temperature Range: Ambient to 300 °C, Temperature Accuracy: ±1 °Overheat Protection: Automatic cut-off with warning indicator, Stirring Speed Range: 100 to 1500 rpm,Maximum Stirring Volume: Up to 10L (water), Power Supply: 230 V, 50 Hz, digital. | 01 |  |  |
| 36 | Grab sampler | 3L Sediment Sampler, Stainless Steel | 01 |  |  |
| 37 | Wooden Herbarium Press | EISCO Wooden Herbarium Press Frame: 12 x 18 inches with 2 Buckle Straps | 01 |  |  |
| 38 | Herbarium sheets (Pack of 25) | Archival Quality, PH 8.5, Acid-Free, Cotton Rag Unbuffered Material for Institutions and Museums-11.5x16.5 inch, 0.015'' Thickness- Pack of 25 | 01 |  |  |
| 39 | Spectrophotometer – Double Beam UV-Vis, 2 nos. Quartz Cuvettes (10 mm), Optical Glass Cuvettes (10 mm) | UV-Vis Spectrophotometer (190–1100 nm), double beam. Spectral bandwidth 1.0 nm or better; wavelength accuracy ±0.3 nm or better; wavelength repeatability ±0.1 nm or better; stray light ≤0.02% (at 220 nm and 340 nm). Photometric range -4.0 to 4.0 Abs or better. Automatic lamp switching (deuterium and tungsten/halogen). Instrument shall include built-in performance verification and support external validation using certified reference materials. Calibration must be traceable to SI units and carried out by a Sri Lanka Accreditation Board accredited laboratory. Sample holder compatible with 10 mm cuvettes (quartz and optical glass). Supply with quartz and optical glass cuvettes. Supplier shall provide IQ/OQ/PQ documentation, wavelength verification certificate, and local service support. Annual maintenance and calibration costs to be clearly indicated. Bid shall itemize costs for instrument, cuvettes, and accessories. 3 years warranty. | 01 |  |  |
| 40 | Spectrophotometer – Double Beam UV-Vis, 2 nos. Quartz Cuvettes (10 mm), Optical Glass Cuvettes (10 mm) | Matched Pair (02 units) of Quartz Cuvettes (10 mm) – Mandatory for UV/Nitrate testing. | 02 |  |  |
| 41 | Spectrophotometer – Double Beam UV-Vis, 2 nos. Quartz Cuvettes (10 mm), Optical Glass Cuvettes (10 mm) | Matched Pair (02 units) of Optical Glass Cuvettes (10 mm) – For Visible/Ammonia testing. | 02 |  |  |
| 42 | Power Protection (Online UPS) | Online (double conversion) UPS, minimum 3 kVA capacity, pure sine wave output, zero transfer time. Input 160–300 V AC, output 230 V ±1%, frequency 50/60 Hz. THD ≤3%. Backup time minimum 15 minutes at 50% load. Equipped with maintenance-free batteries, LCD display, overload, short circuit, surge protection, and automatic bypass. Minimum 1 year warranty. | 01 |  |  |
| 43 | Buchner funnel | Excellent resistance to acids and alkalis. Thermal Stability: Can typically withstand temperatures up to 1000°C Usually glazed inside and out (except the rim) | 01 |  |  |
| 44 | Benchtop pH/Ion Meter | Benchtop pH/Ion Meter with measurement range -2.00 to 20.00 pH, resolution 0.001 pH, and accuracy ±0.002 pH or better. Calibration: Minimum 3–5 point automatic calibration using standard buffer solutions (pH 4.01, 7.00, 10.01 or equivalent). Temperature: Automatic temperature compensation (ATC) with temperature measurement range 0–100°C. Sensor: Glass-body combination pH electrode with ATC probe. Features: Digital display, data storage and/or USB interface with GLP compliance. Calibration certificate: Traceable to international standards through Sri Lanka Accreditation Board accredited laboratory or equivalent. minimum 2-year warranty. | 01 |  |  |
| 45 | Semi-Automated Kjeldahl Distillation System | semi-automated Kjeldahl distillation unit designed for nitrogen/protein determination. 12 positions or higher. TECHNICAL REQUIREMENTS*Semi-Automated Kjeldahl Distillation Unit Measuring range: 0.1–240 mg N or better Distillation time: 3–6 minutes per sample or better Recovery: ≥99.5% Reproducibility: RSD ≤±1% Automatic NaOH addition and controlled steam distillation Operator-controlled operation with adjustable parameters * Digestion Unit 12-position heating block Temperature range up to 450°C or better Accuracy ±1°C or better Compatible with 300 mL tubes * Accessories Borosilicate digestion tubes (300 mL) – 12 Nos Reagent tanks (NaOH, boric acid, waste) – 01 set Tubing, connectors–01set *Exhaust / Scrubber System Basic fume neutralizer: INSTALLATION AND COMMISSIONING Delivery, installation, testing, commissioning and operational demonstration at site. CALIBRATION AND VALIDATION IQ/OQ/PQ documentation Performance verification Calibration traceable to SI units via accredited authority (e.g., SLAB or equivalent) Minimum 2-year warranty. Annual maintenance and calibration cost to be quoted separately. | 01 |  |  |
| 46 | Soxhlet Extraction Apparatus (3-Place unit) | Complete Set (Consisting of an Extraction Body, Condenser, and Boiling Flask.), Boiling flask -High-quality Borosilicate Glass 3.3; high thermal shock resistance and chemical durability. Capacity: 500 ml. Extractor Body- 250 ml nominal capacity; compatible with the 500ml flask., Condenser-Allin (Bulb) or Dimorph type; approx. 250–300mm length. Cellulose Thimbles Pack-Must fit the 250ml extractor (typically approx. 33 X 80mm or 33 X 94 mm).5 meters of high-grade, flexible Silicone or PVC tubing. Must be chemical-resistant and capable of withstanding standard laboratory tap water pressure without ballooning. Clamps, integrated manifold for distributed water flow to all three condensers. Heater (3-Place), support stands and clamps. minimum 2-year warranty. | 01 |  |  |
| 47 | Laboratory Muffle Furnance | Temperature: Ambient to 1100 oC (operating range 550-600 oC for ash) or better. Control: Digital PID programmable controller with +/-1 oC accuracy. Chamber Volume: Minimum 7Liters.Safety: Automatic power-off when the door is opened; ceramic fiber insulation for low outer shell temperature. Bench-top type with compact dimensions, Exhaust pipe at rear wall, with vapor vent connection, Single phase, 230V / 50Hz, minimum 2-year comprehensive warranty. | 01 |  |  |
| 48 | High Performance Fume Hood | High-performance ducted laboratory fume hood, 4 ft width, complete with base cabinet, work surface, service fixtures, electrical provisions, and full exhaust ducting system suitable for chemical laboratory applications.Type: Ducted high-performance chemical fume hood Nominal width: 4 ft (1200 mm approx.) Internal construction: Chemical-resistant, corrosion-proof liner (polypropylene / epoxy-coated steel / phenolic resin or equivalent) External body: Powder-coated steel or equivalent corrosion-resistant material Sash: Vertical sliding, counterbalanced safety glass (tempered) Airflow face velocity: 0.4 – 0.6 m/s Lighting: LED or fluorescent internal lighting system, Work Surface Material: Solid epoxy resin / phenolic resin / ceramic (chemical resistant) Acid, alkali, and heat resistant, Provision for installation of service fixtures on both sides of hood wall Includes mounting points for: Gas, Water, Air lines Chemical-resistant valves (brass/PTFE coated or equivalent), Built-in electrical sockets (minimum 2 nos, GFCI protected or equivalent) Power rating: 220–240 V, 50/60 Hz External electrical outlets allowed on both side walls Integrated cable management system, Exhaust System (Ducting) Type: Ducted exhaust system for external ventilation Material: PVC / FRP / stainless steel or chemical-resistant ducting Duct diameter: Suitable for airflow requirement Airflow capacity: Designed to maintain stable face velocity under full sash opening Includes: Exhaust blower (remote or inline), Transition connectors, Exhaust discharge designed to prevent re-entry of fumes into laboratory environment External shut-off control accessible from front panel. Minimum 3-year warranty. | 01 |  |  |
| 49 | Rutner Sampler | Ruttner Water Sampler (1 L capacity), vertical type, made of transparent acrylic tube with stainless steel/brass frame, messenger-operated closing mechanism, complete with internal non-mercury thermometer (–10 °C to +60 °C), discharge valve, drop messenger, and suspension line. Suitable for sampling at discrete depths in freshwater and marine environments. | 02 |  |  |
| 50 | Heating Mantel | Designed for 500ml Round Bottom Flasks., Temperature controller, temperature Up to 450 0C (approximate).220V, 50/60Hz | 01 |  |  |
| 51 | Laboratory Water Distiller | Fully automatic, wall-mounted or benchtop water still. Minimum 4 liters per hour (L/h) of distilled water. Pyrogen-free distillate; conductivity approximately 3.0–4.0 µS/cm at 25°C. Boiler and condenser made of borosilicate glass (3.3) or high-grade stainless steel for purity and corrosion resistance. Automatic cut-off with low-water level safety switch to prevent dry running. Automatic restart when water supply is restored. Thermal overload protection against overheating. Power supply: 220–240 V, 50/60 Hz. minimum 2-year warranty. | 02 |  |  |
| 52 | Top Loading Balance | Capacity 10kg, readability 0.1g, LCD display, at least 1 year warranty | 01 |  |  |
| 53 | BOD Incubator (Refrigerated) | For Capacity: 12 BOD bottles (300 mL each), maintained 20 0C ±10C, Digital temperature controller Forced air circulation for uniformity. | 01 |  |  |
| 54 | Mufffle Furnace | Max Temp: 1200°C; Working Temp: 1100°C; Chamber Volume: 7L; Heating Element: Fe-Cr-Al alloy; Controller: PID Digital; Safety: Over-temperature protection; Power: 3.5kW. | 01 |  |  |
| 55 | Thermometer (Digital) | Digital probe for precise temperature monitoring in tanks and laboratory experiments. | 03 |  |  |
| 56 | Water Quality Multi Meter | Parameters: pH, DO, Cond, TDS, Salinity, Temp; pH Range: 0-14; DO Range: 0-20 mg/L; Salinity: 0-70 ppt; Display: Backlit LCD; Protection: IP67 waterproof. | 01 |  |  |
| 57 | Centrifuge | Max Speed: 4000 rpm; Max RCF: 2325 xg; Capacity: 12 x 15ml; Timer: 0-99 mins; Safety: Lid lock, overspeed protection; Power: 220V/50Hz. | 01 |  |  |
| 58 | Magnetic Stirrer | Speed Range: 100-1500 rpm; Max Stirring Vol: 5L; Plate Material: Ceramic coated; Power: 220V; Heating: No; Speed Control: Analog dial. | 01 |  |  |
| 59 | pH Meter with Probe | Range: -2.00 to 16.00 pH; Resolution: 0.01 pH; Accuracy: ±0.01 pH; Temp Compensation: Automatic (ATC); Calibration: 3-point automatic; Probe: Glass Body BNC. | 01 |  |  |
| 60 | Dissecting Set (Fish) | Includes scalpel, scissors, forceps, dissecting needle and tray; stainless steel. | 10 |  |  |
| 61 | Dissecting Tray with Wax | Approx. 12” × 8”, wax base, durable frame (plastic/aluminum). | 10 |  |  |
| 62 | Thermometer | Range 0°C to 100°C, alcohol or mercury. | 05 |  |  |
| 63 | Water Quality Testing Meter | Portable, digital, measures pH, DO, temperature, and TDS. | 01 |  |  |
| 64 | Refractometer | Handheld, salinity measurement with temperature compensation. | 02 |  |  |
| 65 | Analytical Balance | Digital, 4 decimal accuracies (0.0001g), calibration feature. | 01 |  |  |